## Slide 1
Sơ đồ nguyên lý liên kết HIL
A1/A2/A3 là mã cụm. Đấu theo tên chân in trên bo, không suy số chân theo vị trí trên hình.
A1 BỘ PHÁT
ESP32 DevKit V1
ESP-WROOM-32
GPIO25 / DAC
GND
Micro-USB 5 V
A2 BỘ THU HIL
TMe ESP32-S3 Dev
WROOM-1 N8R8
GPIO1 / ADC1_CH0
ADC 12-bit, atten cao nhất
GND
USB-C 5 V
3V3
GPIO8 / SDA
GPIO9 / SCL
GND
A3 OLED
SSD1306, I²C
VCC / 3V3
SDA
SCL
GND
R1 10 kΩ ±5%
DAC_OUT
ADC_IN
R2 39 kΩ ±5%
C1 100 nF ±20%
50 V
sát ADC
TP1
TP2
TP3
GND / 0 V
3V3_A2
R3 4k7
DNP*
R4 4k7
DNP*
3V3_A2
C2
100 nF / 50 V
sát A3
GND
1. Chấm đen = nối. Giao dây không chấm = không nối. GND là mass tín hiệu.
2. R3/R4: DNP nếu A3 đã có pull-up phù hợp về 3V3_A2. Nếu thiếu, lắp 4k7.
3. I²C 100 kHz. C1 sát GPIO1, C2 sát OLED. Hai bo cấp USB riêng, không nối 5V.
4. Đo TP1/TP2 so với TP3. TP2 ≤2,85 V khi nguồn A1 nằm trong 3,3 V ±5%.
RF-HIL-001 / Liên kết nguyên mẫu HIL
Rev B / 18-10-2026
Tờ 1 / 1
Duyệt: chưa phê duyệt
NGUYÊN MẪU HIL / Chưa phát hành sản xuất / Bản vẽ theo giao diện module, không phải layout PCB
Notes: RF-HIL-001 Rev B. Module-interface schematic only, not a PCB production release.
Drawing conventions informed by IEC 61082-1:2014 public scope and IEC 60617. No clause-by-clause conformance certification.
https://webstore.iec.ch/en/publication/4469
https://tc3.iec.ch/tc-activity/graphical-symbols-for-diagrams/
GPIO25 DAC: https://docs.espressif.com/projects/esp-idf/en/latest/esp32/api-reference/peripherals/dac.html
GPIO1 ADC1_CH0 and 100 nF: https://docs.espressif.com/projects/esp-hardware-design-guidelines/en/latest/esp32s3/schematic-checklist.html
ADC accuracy guidance states 0..2900 mV at highest attenuation; Arduino measurable range is broader (0..3100 mV). This drawing targets the narrower range.
I2C: https://www.nxp.com/docs/en/user-guide/UM10204.pdf and https://www.ti.com/lit/an/slva689/slva689.pdf
A2 intended purchase is TMe ESP32-S3 Dev (https://www.tme.vn/Products.aspx?CateId=457), not Espressif DevKitC-1. TMe individual product/PDF could not be re-fetched. Verify actual board revision and silk labels before assembly.
OLED pin order and installed pull-ups are not frozen. R3/R4 are optional 4k7 to 3V3_A2. DNP when existing module pull-ups meet the effective resistance and rise-time acceptance conditions. C2 local decoupling is added.
R2 changed 56k to 39k. For ideal DAC output bounded by 3.3 V +5% and resistor tolerances +/-5%, worst ADC_IN is below 2.85 V. Actual DAC source impedance and S3 ADC transfer curve are not modeled.
R1/R2 budget and module prices remain previous estimates. New 39k procurement and optional 4k7 require quotation. Hardware fit and production approval are not complete.

## Slide 2
Phần cứng cần mua
01 bộ tối thiểu, giá chưa gồm vận chuyển
RELYFETAL / HIL
CẤU HÌNH REV B
R3/R4 dự phòng chưa tính giá. Chốt revision, nhãn chân và pull-up của A1–A3 trước khi đặt hàng.
02 / 03
BỘ CƠ BẢN, CHƯA GỒM R3/R4
≈622.000₫
SL | Linh kiện / cấu hình | Giá dự kiến
1 | A1 ESP32 DevKit V1 ESP-WROOM-32, Micro-USB | 109.000₫
1 | A2 TMe ESP32-S3 Dev WROOM-1 N8R8, USB-C | 250.000₫
1 | A3 OLED SSD1306 0,96 in I²C, VCC 3,3 V, kiểm tra nhãn chân | 60.086₫
2 gói+ | R1 10 kΩ + R2 39 kΩ, 1/4 W, ±5% R3/R4 4k7 dự phòng, DNP mặc định | ≈5.500₫ + R3/R4
3 món | C1/C2 100 nF/50 V + MB-102 830 lỗ Dây Dupont 40P, tụ lấy cùng gói 10 chiếc | ≈69.900₫
2 | Cáp dữ liệu Micro-USB và USB-A sang USB-C Mỗi loại 1 dây, dài 1 m | ≈127.300₫
Notes: Danh sách đã đồng bộ với sơ đồ slide 1: R1=10 kΩ, R2=56 kΩ, C1=100 nF; cầu chia áp đưa tối đa 3,3 V xuống khoảng 2,8 V. N8R8 official board has 8 MB flash + 8 MB PSRAM https://docs.espressif.com/projects/esp-dev-kits/en/latest/esp32s3/esp32-s3-devkitc-1/user_guide_v1.1.html ; GPIO1 is ADC1_CH0 and ADC 12 dB nominal range is about 0–3.1 V https://docs.espressif.com/projects/esp-idf/en/latest/esp32s3/api-reference/peripherals/adc.html . Clone-board pin mapping can vary; verify silkscreen/module before ordering.
KIỂM THỬ MÔ PHỎNG (18/10/2026)
Chạy python3 analysis/hardware_simulator/simulate_hil_rc.py. Mô hình dùng ECG_RECORD_R01/A02 (2.500 mẫu/ca ở 250 Hz), DAC 0–3,3 V lý tưởng, R1=10 kΩ, R2=56 kΩ, C1=100 nF, 12-bit ADC lý tưởng và lệch pha nguồn/thu 0,1–3,9 ms. Kết quả: gain=0,8485; tau=0,848 ms; fc=187,6 Hz; VADC peak danh định 2,800 V, góc R±5% cao nhất 2,841 V. Ở pha 0,1 ms, tương quan dạng sóng với DAC lý tưởng là 0,876 (r01) và 0,754 (a02); khi pha 3,9 ms gần 1. Suy ra điện áp không vượt 3,0 V mục tiêu trong mô hình, nhưng firmware lấy mẫu tự do 250 Hz không đồng bộ, nên kết quả dạng sóng phụ thuộc pha. ADC mean lý tưởng ca a02 ≈2.094 mã, dưới ngưỡng 2.150 của gate hiện tại; gate demo có thể không từ chối ca a02. Đây là mô phỏng tín hiệu và kiểm tra mã, chưa thay thế đo oscilloscope/ADC thật. Firmware hiện tại cũng chưa nạp/chạy TCN thật.
GIÁ TRA CỨU (VNĐ, chưa gồm vận chuyển; giá/biến thể có thể đổi)
J1 Shopee, ESP32 DevKit V1 ESP-WROOM-32, 109.000: https://shopee.vn/ESP32-DevKit-V1-ESP32-WROOM-32-WiFi-Bluetooth-Development-Board-%E2%80%93-Kit-Ph%C3%A1t-Tri%E1%BB%83n-IoT-Ch%C3%ADnh-H%C3%A3ng-i.65776741.42482058047
J2 TMe, ESP32-S3 Dev dựa ESP32-S3-WROOM-1 N8R8, mã trang 2504, 250.000: https://tme.vn/Product.aspx?id=2504
J3 Shopee PCBLab, SSD1306 0.96in 128x64 4-pin, 60.086: https://shopee.vn/M%C3%A0n-h%C3%ACnh-OLED-0.96-inch-I2C-SSD1306-128x64-giao-ti%E1%BA%BFp-I2C-4-ch%C3%A2n-VCC-GND-SCL-SDA-d%C3%B9ng-cho-Arduino-ESP32-STM32-i.1723076079.51514172829
R1 10 kΩ and R2 56 kΩ, carbon film through-hole 1/4 W 5%, Thegioiic 55 each, MOQ 50 each; R2 seller code IC929: https://www.thegioiic.com/dien-tro-56-kohm-1-4w-5-4-vong-mau ; 10 kΩ listing: https://www.thegioiic.com/dien-tro-10-kohm-1-4w-5-4-vong-mau
C1 ceramic 104 100 nF 50 V, Shopee 10 pcs 6.000: https://shopee.vn/T%E1%BB%A5-g%E1%BB%91m-104-100nF-50V-Ceramic-%2810pcs%29-i.29186330.43423148258
Breadboard MB-102 830 tie points, 32.000: https://shopee.vn/Breadboard-MB-102-830-l%E1%BB%97-Nh%E1%BB%B1a-ABS-ch%E1%BA%A5t-l%C6%B0%E1%BB%A3ng-cao-i.120349161.5216095760 ; Dupont 40P, 21 cm, 31.900–39.900: https://shopee.vn/-B%E1%BB%99-S%E1%BB%A3i-D%C3%A2y-N%E1%BB%91i-40P-D%C3%A0i-21Cm-%28-2-%C4%90%E1%BA%A7u-%C4%90%E1%BB%B1c-C%C3%A1i-%29-D%C3%A2y-C%E1%BA%AFm-Test-M%E1%BA%A1ch--i.61435118.25088627856
USB Micro-B data cable 1 m, Shopee 47.880: https://shopee.vn/C%C3%A1p-d%E1%BB%AF-li%E1%BB%87u-s%E1%BA%A1c-nhanh-USB-Micro-USB-1m-i.376801883.28332691601 ; USB-A to USB-C data cable 1 m, USAMS Shopee 79.420: https://shopee.vn/Usams-D%C3%A2y-C%C3%A1p-S%E1%BA%A1c-Nhanh-Truy%E1%BB%81n-D%E1%BB%AF-Li%E1%BB%87u-%C4%90%E1%BA%A7u-USB-Type-C-D%C3%A0i-1M-D%C3%A0nh-Cho-%C4%90i%E1%BB%87n-Tho%E1%BA%A1i-i.408923716.22343546073
ESP32 DAC GPIO25 official source: https://docs.espressif.com/projects/esp-idf/en/latest/esp32/api-reference/peripherals/dac.html ; ESP32-S3 GPIO1 ADC1_CH0 and calibration caution: https://docs.espressif.com/projects/esp-hardware-design-guidelines/en/latest/esp32s3/schematic-checklist.html ; N8R8 ordering code and board USB supply options: https://documentation.espressif.com/esp-dev-kits/en/latest/esp32s3/esp32-s3-devkitc-1/user_guide_v1.1.html
*Total is approximately 622.000 VND including minimum resistor packs; excludes shipping and marketplace promotions.
REV B UPDATE:
RF-HIL-001 Rev B. Module-interface schematic only, not a PCB production release.
Drawing conventions informed by IEC 61082-1:2014 public scope and IEC 60617. No clause-by-clause conformance certification.
https://webstore.iec.ch/en/publication/4469
https://tc3.iec.ch/tc-activity/graphical-symbols-for-diagrams/
GPIO25 DAC: https://docs.espressif.com/projects/esp-idf/en/latest/esp32/api-reference/peripherals/dac.html
GPIO1 ADC1_CH0 and 100 nF: https://docs.espressif.com/projects/esp-hardware-design-guidelines/en/latest/esp32s3/schematic-checklist.html
ADC accuracy guidance states 0..2900 mV at highest attenuation; Arduino measurable range is broader (0..3100 mV). This drawing targets the narrower range.
I2C: https://www.nxp.com/docs/en/user-guide/UM10204.pdf and https://www.ti.com/lit/an/slva689/slva689.pdf
A2 intended purchase is TMe ESP32-S3 Dev (https://www.tme.vn/Products.aspx?CateId=457), not Espressif DevKitC-1. TMe individual product/PDF could not be re-fetched. Verify actual board revision and silk labels before assembly.
OLED pin order and installed pull-ups are not frozen. R3/R4 are optional 4k7 to 3V3_A2. DNP when existing module pull-ups meet the effective resistance and rise-time acceptance conditions. C2 local decoupling is added.
R2 changed 56k to 39k. For ideal DAC output bounded by 3.3 V +5% and resistor tolerances +/-5%, worst ADC_IN is below 2.85 V. Actual DAC source impedance and S3 ADC transfer curve are not modeled.
R1/R2 budget and module prices remain previous estimates. New 39k procurement and optional 4k7 require quotation. Hardware fit and production approval are not complete.

## Slide 3
Kiểm thử HIL trên phần cứng
RELYFETAL / HIL
01
Chốt mã bo & firmware
Xác nhận revision A1–A3, ADC attenuation và timer lấy mẫu.
02
Nghiệm thu điện & I²C
TP2 ≤2,85 V, chung GND. I²C 100 kHz, quét địa chỉ OLED.
03
Nạp pipeline thật
Thu 1 kS/s trong cửa sổ 4 s; chạy tiền xử lý và TCN INT8 trên bo.
04
So khớp hai ca
r01: đối chiếu F1 và jitter. a02: kiểm tra gate, ẩn fHR.
Lưu raw ADC và log.
TIÊU CHÍ HIL
Lấy mẫu 1 kHz
Không mất mẫu, có log jitter
ADC không clipping
Đo TP2 và hiệu chuẩn ADC
TCN chạy trên bo
Khớp offline theo dung sai
Hiện tại: firmware demo 250 Hz, chưa TCN trên bo.
Rev B: nguyên mẫu. Cần chốt module, ERC ECAD và đo bench trước phát hành sản xuất.
03 / 03
Notes: Chi tiết đối chiếu repo: analysis/KIENTRUC.md ghi pipeline huấn luyện/chấm ở 1.000 Hz và phép biến đổi xuống 250 Hz; analysis/hardware_simulator/esp32_patient_simulator.ino hiện phát 250 Hz bằng micros() và in Serial mỗi mẫu; esp32s3_receiver_oled.ino đặt current_fhr = 142, quyết định safe theo trung bình ADC, không gọi mô hình. Vì vậy không được mô tả kết quả firmware demo là kết quả HIL của TCN. Với ADC, 12 dB có dải đo danh định 0–3,1 V; divider giảm đầu ra DAC để có biên an toàn: https://docs.espressif.com/projects/esp-idf/en/latest/esp32s3/api-reference/peripherals/adc.html .
REV B UPDATE:
RF-HIL-001 Rev B. Module-interface schematic only, not a PCB production release.
Drawing conventions informed by IEC 61082-1:2014 public scope and IEC 60617. No clause-by-clause conformance certification.
https://webstore.iec.ch/en/publication/4469
https://tc3.iec.ch/tc-activity/graphical-symbols-for-diagrams/
GPIO25 DAC: https://docs.espressif.com/projects/esp-idf/en/latest/esp32/api-reference/peripherals/dac.html
GPIO1 ADC1_CH0 and 100 nF: https://docs.espressif.com/projects/esp-hardware-design-guidelines/en/latest/esp32s3/schematic-checklist.html
ADC accuracy guidance states 0..2900 mV at highest attenuation; Arduino measurable range is broader (0..3100 mV). This drawing targets the narrower range.
I2C: https://www.nxp.com/docs/en/user-guide/UM10204.pdf and https://www.ti.com/lit/an/slva689/slva689.pdf
A2 intended purchase is TMe ESP32-S3 Dev (https://www.tme.vn/Products.aspx?CateId=457), not Espressif DevKitC-1. TMe individual product/PDF could not be re-fetched. Verify actual board revision and silk labels before assembly.
OLED pin order and installed pull-ups are not frozen. R3/R4 are optional 4k7 to 3V3_A2. DNP when existing module pull-ups meet the effective resistance and rise-time acceptance conditions. C2 local decoupling is added.
R2 changed 56k to 39k. For ideal DAC output bounded by 3.3 V +5% and resistor tolerances +/-5%, worst ADC_IN is below 2.85 V. Actual DAC source impedance and S3 ADC transfer curve are not modeled.
R1/R2 budget and module prices remain previous estimates. New 39k procurement and optional 4k7 require quotation. Hardware fit and production approval are not complete.